## Instructions
Instructions for Analyzing QIAseq DNA QuantiMIZE 384-Well Array/Assay Kit Results for the QIAseq Targeted DNA Panels with this file:
Generally, only change data in yellow cells. Gray and white cells contain formulas for calculation or results. Please do not change them.
1. Assay Quick Start Guide: QuantiMIZE Assay Kit Guide - Input your number of Samples into the yellow "enter number" Cell D7, and then follow the protocol step-by-step to set up assays with your Samples' gDNA and the Control gDNA.
2. Array Quick Start Guide: QuantiMIZE Array Kit Guide - Follow the protocol step-by-step to set up the array with your Samples' gDNA and the Control gDNA.
3. Raw Data & Analysis Setup Worksheet: a. Copy and Paste Special Values the raw CT values from the real-time PCR instrument run of your Samples' gDNA and the Control gDNA into the yellow section under the "Raw Data" column, Cells B2 through B385. When a "Delete the Outlier" message appears in the "Message" column F, identify the obvious outlier CT value in Column B and manually delete it to insure the "Message" disappears. b. Input the Samples' names into the yellow Cells I2 through I64. Each Sample must have a unique name. c. Input the Samples' concentrations (as measured by NanoDrop or other methods) into the yellow Cells J2 through J64. d. Delete any "Raw Data", "Sample ID", and "Sample Conc" information provided as an example that is not yours.
4. Results: a. Quality Control Results: Assay QC: If the "Assay QC" for all the Samples' gDNA and the Control gDNA reports “Pass”, then the assays performed optimally. If the Control gDNA "Assay QC" reports “Fail”, double check the assay setup and real-time PCR program. If any Sample’s "Assay QC" reports "Fail", the Sample may contain PCR inhibitor contaminants. b. Quality Control Results: QC Score & QC Call: The "QC Score" provides an indication of DNA Sample damage or fragmentation. Samples with a "QC Score" <= 0.04 are given a "High" "QC Call", meaning high quality DNA and a low degree of damage or fragmentation. Samples with a "QC Score" > 0.04 are given a "Low" "QC Call", meaning low quality DNA and a high degree damage or fragmentation. Samples with a "Low" "QC Call" may only be suitable for detection of high-frequency variants. c. Sample Volume and Amount Recommended: Columns F and G list the "Sample Volume" and "Sample Amount” ranges that are recommended for each Sample. If possible, the highest amount of Samples should be used. d. Comments: Even if a Sample is noted with the "Comment" "We do not recommend proceeding with this Sample.", up to 500 ng of Sample gDNA may still be used if a Sample must be processed.
Version 1.0, February 2017

## Assay Quick Start Guide
Quick Start Guide for QIAseq DNA QuantiMIZE Assay Kit
Always include QuantiMIZE Control gDNA in every run as described below, regardless of the number of gDNA Samples processed.
Each gDNA Sample is interrogated by 6 reactions: 3 (triplicate) reactions using Assay100 and 3 (triplicate) reactions using Assay200.
1. Prepare 2 PCR Components Mixes, Mix 1 using Primer Assay100 and Mix 2 for Primer Assay200. For a complete recipe, input number of Samples to be analyzed into the yellow "enter number" Cell D7.
Component | Volume (µL) | Sample number (n)
RNase-/DNase-free water | enter number in yellow box | enter number
QIAseq qPCR SYBR Green Mastermix | enter number in yellow box
Primer Assay100 or Assay200 | enter number in yellow box
Total volume | enter number in yellow box
2. Dispense 8 μL of PCR Components Mix 1 (Assay100) or Mix 2 (Assay200) into the appropriate Sample-specific (up to the number of Samples being processed) and Control white or gray wells, respectively, according to layout below:
3. Dilute each Sample gDNA or Control gDNA separately with RNase-/DNase-free water according to the Table below.
Component | Volume
Genomic DNA (Control or Sample) | 2 μL
RNase-/DNase-free water | 14 μL
Total volume | 16 μL
4. Dispense 2μL of each diluted Sample gDNA or Control gDNA into the appropriate wells according to the Layout described in step 2 above. Sample 1 (S1) occupies wells A01 through A03 for Assay100 (white cells) and A13 through A15 for Assay200 (shaded cells). Sample 2 (S2) occupies wells B01 through B03 and B13 through B15, and so on up to the number of Samples being processed. The diluted Control gDNA MUST BE ADDED to Wells P10 through P12 for Assay100 and P22 through P24 for Assay200.
5. Seal the plate, vortex and spin it down.
6. Run the following real-time PCR cycling program:

## Array Quick Start Guide
Quick Start Guide for QIAseq DNA QuantiMIZE Array Kit
Always include QuantiMIZE Control gDNA in every run as described below, regardless of the number of gDNA Samples processed.
Each gDNA Sample is interrogated by 6 reactions: 3 (triplicate) reactions using Assay100 and 3 using (triplicate) reactions Assay200.
1. Prepare a PCR Components Mix for each Sample gDNA and Control gDNA using the following recipe:
Component | For each Sample (6 wells)
RNase-/DNase-free water | 34 µl
QIAseq qPCR SYBR Green Mastermix | 36 µl
Genomic DNA (Control or Sample) | 2 µl
Final volume | 72 µl
2. Dispense 10 μL of each Sample-specific (up to the number of Samples being processed) and Control PCR Components Mix into the appropriate wells indicated by the Array Layout below.
Add the Sample 1 (S1) PCR Components Mix to wells A01 through A03 for Assay100 (white cells) and A13 through A15 for Assay200 (shaded cells). Add the Sample 2 (S2) PCR Components Mix to wells B01 through B03 and B13 through B15, and so on up to the number of Samples being processed. The Control gDNA PCR Components Mix MUST BE ADDED to Wells P10 through P12 for Assay100 and P22 through P24 for Assay200.
3. Seal the plate, vortex and spin it down.
4. Run the following real-time PCR cycling program:

## Raw Data & Analysis Setup
Well | Raw Data | Average CT | Standard Deviation | Message | Sample | Sample ID | Sample Conc. (ng/µL)
A01 | 23.328 | 23.351 | 0.039837 |  | S1: | Thyroid Tumor | 36 | Please input sample concentrations for calculations! Use NanoDrop or other methods.
A02 | 23.397 |  |  |  | S2: | Normal Thyroid | 23
A03 | 23.328 |  |  |  | S3: | Melanoma 1 | 27
A04 | 19.319 | 19.3193 | 0.012504 |  | S4: | Paired Breast Tumor | 13
A05 | 19.332 |  |  |  | S5: | Paired Normal Breast | 9
A06 | 19.307 |  |  |  | S6: | Melanoma 2 | 31
A07 | 18.198 | 18.1893 | 0.029956 |  | S7: | NTS (all-prep) | 22
A08 | 18.156 |  |  |  | S8: | N30 (all-prep) | 52
A09 | 18.214 |  |  |  | S9: | PLT | 14
A10 | 25.905 | 25.8487 | 0.058621 |  | S10: | PLN | 5
A11 | 25.788 |  |  |  | S11: | FFPE Melanoma | 11
A12 | 25.853 |  |  |  | S12: | FFPE Colon | 50
A13 | 30.212 | 29.9747 | 0.20566 |  | S13: | N19 | 14
A14 | 29.863 |  |  |  | S14: | N10 | 20
A15 | 29.849 |  |  |  | S15: | NA12878 | 20
A16 | 22.251 | 22.2453 | 0.0055074 |  | S16: | Thyroid Tumor | 36
A17 | 22.24 |  |  |  | S17: | Normal Thyroid | 23
A18 | 22.245 |  |  |  | S18: | Melanoma 1 | 27
A19 | 20.559 | 20.585 | 0.032187 |  | S19: | Paired Breast Tumor | 13
A20 | 20.575 |  |  |  | S20: | Paired Normal Breast | 9
A21 | 20.621 |  |  |  | S21: | Melanoma 2 | 31
A22 |  | 31.4855 | 0.16758 |  | S22: | NTS (all-prep) | 22
A23 | 31.604 |  |  |  | S23: | N30 (all-prep) | 52
A24 | 31.367 |  |  |  | S24: | PLT | 14
B01 | 19.319 | 19.3193 | 0.012504 |  | S25: | PLN | 5
B02 | 19.332 |  |  |  | S26: | FFPE Melanoma | 11
B03 | 19.307 |  |  |  | S27: | FFPE Colon | 50
B04 | 18.198 | 18.1893 | 0.029956 |  | S28: | N19 | 14
B05 | 18.156 |  |  |  | S29: | N10 | 20
B06 | 18.214 |  |  |  | S30: | NA12878 | 20
B07 | 25.905 | 25.8487 | 0.058621 |  | S31: | Thyroid Tumor | 36
B08 | 25.788 |  |  |  | S32: | Normal Thyroid | 23
B09 | 25.853 |  |  |  | S33: | Melanoma 1 | 27
B10 | 24.839 | 24.8837 | 0.040808 |  | S34: | Paired Breast Tumor | 13
B11 | 24.919 |  |  |  | S35: | Paired Normal Breast | 9
B12 | 24.893 |  |  |  | S36: | Melanoma 2 | 31
B13 | 22.251 | 22.2453 | 0.0055074 |  | S37: | NTS (all-prep) | 22
B14 | 22.24 |  |  |  | S38: | N30 (all-prep) | 52
B15 | 22.245 |  |  |  | S39: | PLT | 14
B16 | 20.559 | 20.585 | 0.032187 |  | S40: | PLN | 5
B17 | 20.575 |  |  |  | S41: | FFPE Melanoma | 11
B18 | 20.621 |  |  |  | S42: | FFPE Colon | 50
B19 |  | 31.4855 | 0.16758 |  | S43: | N19 | 14
B20 | 31.604 |  |  |  | S44: | N10 | 20
B21 | 31.367 |  |  |  | S45: | NA12878 | 20
B22 | 29.804 | 29.7497 | 0.28146 |  | S46: | Thyroid Tumor | 36
B23 | 29.445 |  |  |  | S47: | Normal Thyroid | 23
B24 | 30 |  |  |  | S48: | Melanoma 1 | 27
C01 | 18.198 | 18.1893 | 0.029956 |  | S49: | Paired Breast Tumor | 13
C02 | 18.156 |  |  |  | S50: | Paired Normal Breast | 9
C03 | 18.214 |  |  |  | S51: | Melanoma 2 | 31
C04 | 25.905 | 25.8487 | 0.058621 |  | S52: | NTS (all-prep) | 22
C05 | 25.788 |  |  |  | S53: | N30 (all-prep) | 52
C06 | 25.853 |  |  |  | S54: | PLT | 14
C07 | 24.839 | 24.8837 | 0.040808 |  | S55: | PLN | 5
C08 | 24.919 |  |  |  | S56: | FFPE Melanoma | 11
C09 | 24.893 |  |  |  | S57: | FFPE Colon | 50
C10 | 18.062 | 18.0493 | 0.035725 |  | S58: | N19 | 14
C11 | 18.077 |  |  |  | S59: | N10 | 20
C12 | 18.009 |  |  |  | S60: | NA12878 | 20
C13 | 20.559 | 20.585 | 0.032187 |  | S61: | Thyroid Tumor | 36
C14 | 20.575 |  |  |  | S62: | Normal Thyroid | 23
C15 | 20.621 |  |  |  | S63: | Melanoma 1 | 27
C16 |  | 31.4855 | 0.16758 |  | Control: | Control | Control
C17 | 31.604
C18 | 31.367
C19 | 29.804 | 29.7497 | 0.28146
C20 | 29.445
C21 | 30
C22 | 20.367 | 20.3807 | 0.031786
C23 | 20.417
C24 | 20.358
D01 | 25.905 | 25.8487 | 0.058621
D02 | 25.788
D03 | 25.853
D04 | 24.839 | 24.8837 | 0.040808
D05 | 24.919
D06 | 24.893
D07 | 18.062 | 18.0493 | 0.035725
D08 | 18.077
D09 | 18.009
D10 | 17.841 | 17.7783 | 0.055374
D11 | 17.736
D12 | 17.758
D13 |  | 31.4855 | 0.16758
D14 | 31.604
D15 | 31.367
D16 | 29.804 | 29.7497 | 0.28146
D17 | 29.445
D18 | 30
D19 | 20.367 | 20.3807 | 0.031786
D20 | 20.417
D21 | 20.358
D22 | 19.567 | 19.5553 | 0.015308
D23 | 19.561
D24 | 19.538
E01 | 24.839 | 24.8837 | 0.040808
E02 | 24.919
E03 | 24.893
E04 | 18.062 | 18.0493 | 0.035725
E05 | 18.077
E06 | 18.009
E07 | 17.841 | 17.7783 | 0.055374
E08 | 17.736
E09 | 17.758
E10 | 16.369 | 16.3487 | 0.01877
E11 | 16.345
E12 | 16.332
E13 | 29.804 | 29.7497 | 0.28146
E14 | 29.445
E15 | 30
E16 | 20.367 | 20.3807 | 0.031786
E17 | 20.417
E18 | 20.358
E19 | 19.567 | 19.5553 | 0.015308
E20 | 19.561
E21 | 19.538
E22 | 17.655 | 17.64 | 0.014526
E23 | 17.626
E24 | 17.639
F01 | 18.062 | 18.0493 | 0.035725
F02 | 18.077
F03 | 18.009
F04 | 17.841 | 17.7783 | 0.055374
F05 | 17.736
F06 | 17.758
F07 | 16.369 | 16.3487 | 0.01877
F08 | 16.345
F09 | 16.332
F10 | 19.644 | 19.6443 | 0.013503
F11 | 19.631
F12 | 19.658
F13 | 20.367 | 20.3807 | 0.031786
F14 | 20.417
F15 | 20.358
F16 | 19.567 | 19.5553 | 0.015308
F17 | 19.561
F18 | 19.538
F19 | 17.655 | 17.64 | 0.014526
F20 | 17.626
F21 | 17.639
F22 | 22.575 | 22.569 | 0.0060005
F23 | 22.563
F24 | 22.569
G01 | 17.841 | 17.7783 | 0.055374
G02 | 17.736
G03 | 17.758
G04 | 16.369 | 16.3487 | 0.01877
G05 | 16.345
G06 | 16.332
G07 | 19.644 | 19.6443 | 0.013503
G08 | 19.631
G09 | 19.658
G10 | 21.098 | 21.032 | 0.057472
G11 | 20.993
G12 | 21.005
G13 | 19.567 | 19.5553 | 0.015308
G14 | 19.561
G15 | 19.538
G16 | 17.655 | 17.64 | 0.014526
G17 | 17.626
G18 | 17.639
G19 | 22.575 | 22.569 | 0.0060005
G20 | 22.563
G21 | 22.569
G22 | 23.975 | 24.0017 | 0.074662
G23 | 23.944
G24 | 24.086
H01 | 16.369 | 16.3487 | 0.01877
H02 | 16.345
H03 | 16.332
H04 | 19.644 | 19.6443 | 0.013503
H05 | 19.631
H06 | 19.658
H07 | 21.098 | 21.032 | 0.057472
H08 | 20.993
H09 | 21.005
H10 | 21.302 | 21.276 | 0.027056
H11 | 21.278
H12 | 21.248
H13 | 17.655 | 17.64 | 0.014526
H14 | 17.626
H15 | 17.639
H16 | 22.575 | 22.569 | 0.0060005
H17 | 22.563
H18 | 22.569
H19 | 23.975 | 24.0017 | 0.074662
H20 | 23.944
H21 | 24.086
H22 | 25.154 | 25.2237 | 0.06853
H23 | 25.291
H24 | 25.226
I01 | 19.644 | 19.6443 | 0.013503
I02 | 19.631
I03 | 19.658
I04 | 21.098 | 21.032 | 0.057472
I05 | 20.993
I06 | 21.005
I07 | 21.302 | 21.276 | 0.027056
I08 | 21.278
I09 | 21.248
I10 | 17.584 | 17.5397 | 0.039323
I11 | 17.526
I12 | 17.509
I13 | 22.575 | 22.569 | 0.0060005
I14 | 22.563
I15 | 22.569
I16 | 23.975 | 24.0017 | 0.074662
I17 | 23.944
I18 | 24.086
I19 | 25.154 | 25.2237 | 0.06853
I20 | 25.291
I21 | 25.226
I22 | 20.004 | 20.086 | 0.071359
I23 | 20.134
I24 | 20.12
J01 | 21.098 | 21.032 | 0.057472
J02 | 20.993
J03 | 21.005
J04 | 21.302 | 21.276 | 0.027056
J05 | 21.278
J06 | 21.248
J07 | 17.584 | 17.5397 | 0.039323
J08 | 17.526
J09 | 17.509
J10 | 18.892 | 18.8573 | 0.043155
J11 | 18.871
J12 | 18.809
J13 | 23.975 | 24.0017 | 0.074662
J14 | 23.944
J15 | 24.086
J16 | 25.154 | 25.2237 | 0.06853
J17 | 25.291
J18 | 25.226
J19 | 20.004 | 20.086 | 0.071359
J20 | 20.134
J21 | 20.12
J22 | 21.336 | 21.3533 | 0.025813
J23 | 21.383
J24 | 21.341
K01 | 21.302 | 21.276 | 0.027056
K02 | 21.278
K03 | 21.248
K04 | 17.584 | 17.5397 | 0.039323
K05 | 17.526
K06 | 17.509
K07 | 18.892 | 18.8573 | 0.043155
K08 | 18.871
K09 | 18.809
K10 | 18.864 | 18.845 | 0.016822
K11 | 18.832
K12 | 18.839
K13 | 25.154 | 25.2237 | 0.06853
K14 | 25.291
K15 | 25.226
K16 | 20.004 | 20.086 | 0.071359
K17 | 20.134
K18 | 20.12
K19 | 21.336 | 21.3533 | 0.025813
K20 | 21.383
K21 | 21.341
K22 | 21.26 | 21.293 | 0.029816
K23 | 21.301
K24 | 21.318
L01 | 17.584 | 17.5397 | 0.039323
L02 | 17.526
L03 | 17.509
L04 | 18.892 | 18.8573 | 0.043155
L05 | 18.871
L06 | 18.809
L07 | 18.864 | 18.845 | 0.016822
L08 | 18.832
L09 | 18.839
L10 | 16.68 | 16.6683 | 0.012583
L11 | 16.67
L12 | 16.655
L13 | 20.004 | 20.086 | 0.071359
L14 | 20.134
L15 | 20.12
L16 | 21.336 | 21.3533 | 0.025813
L17 | 21.383
L18 | 21.341
L19 | 21.26 | 21.293 | 0.029816
L20 | 21.301
L21 | 21.318
L22 | 17.135 | 17.1427 | 0.018717
L23 | 17.129
L24 | 17.164
M01 | 18.892 | 18.8573 | 0.043155
M02 | 18.871
M03 | 18.809
M04 | 18.864 | 18.845 | 0.016822
M05 | 18.832
M06 | 18.839
M07 | 16.68 | 16.6683 | 0.012583
M08 | 16.67
M09 | 16.655
M10 | 23.328 | 23.351 | 0.039837
M11 | 23.397
M12 | 23.328
M13 | 21.336 | 21.3533 | 0.025813
M14 | 21.383
M15 | 21.341
M16 | 21.26 | 21.293 | 0.029816
M17 | 21.301
M18 | 21.318
M19 | 17.135 | 17.1427 | 0.018717
M20 | 17.129
M21 | 17.164
M22 | 30.212 | 29.9747 | 0.20566
M23 | 29.863
M24 | 29.849
N01 | 18.864 | 18.845 | 0.016822
N02 | 18.832
N03 | 18.839
N04 | 16.68 | 16.6683 | 0.012583
N05 | 16.67
N06 | 16.655
N07 | 23.328 | 23.351 | 0.039837
N08 | 23.397
N09 | 23.328
N10 | 19.319 | 19.3193 | 0.012504
N11 | 19.332
N12 | 19.307
N13 | 21.26 | 21.293 | 0.029816
N14 | 21.301
N15 | 21.318
N16 | 17.135 | 17.1427 | 0.018717
N17 | 17.129
N18 | 17.164
N19 | 30.212 | 29.9747 | 0.20566
N20 | 29.863
N21 | 29.849
N22 | 22.251 | 22.2453 | 0.0055074
N23 | 22.24
N24 | 22.245
O01 | 16.68 | 16.6683 | 0.012583
O02 | 16.67
O03 | 16.655
O04 | 23.328 | 23.351 | 0.039837
O05 | 23.397
O06 | 23.328
O07 | 19.319 | 19.3193 | 0.012504
O08 | 19.332
O09 | 19.307
O10 | 18.198 | 18.1893 | 0.029956
O11 | 18.156
O12 | 18.214
O13 | 17.135 | 17.1427 | 0.018717
O14 | 17.129
O15 | 17.164
O16 | 30.212 | 29.9747 | 0.20566
O17 | 29.863
O18 | 29.849
O19 | 22.251 | 22.2453 | 0.0055074
O20 | 22.24
O21 | 22.245
O22 | 20.559 | 20.585 | 0.032187
O23 | 20.575
O24 | 20.621
P01 | 23.328 | 23.351 | 0.039837
P02 | 23.397
P03 | 23.328
P04 | 19.319 | 19.3193 | 0.012504
P05 | 19.332
P06 | 19.307
P07 | 18.198 | 18.1893 | 0.029956
P08 | 18.156
P09 | 18.214
P10 | 18.626 | 18.6033 | 0.019857
P11 | 18.589
P12 | 18.595
P13 | 30.212 | 29.9747 | 0.20566
P14 | 29.863
P15 | 29.849
P16 | 22.251 | 22.2453 | 0.0055074
P17 | 22.24
P18 | 22.245
P19 | 20.559 | 20.585 | 0.032187
P20 | 20.575
P21 | 20.621
P22 | 19.169 | 19.196 | 0.02488
P23 | 19.201
P24 | 19.218

## Results
Sample ID | Quality Control Results |  |  | Sample Concentration (ng/µL) | Sample Volume (µL) Recommended | Sample Amount (ng) Recommended | Comments
| Assay QC | QC Score | QC Call
Thyroid Tumor | Pass | 0.06031 | Low | 36 | 3.3 - 6.9 µL | 120-250ng
Normal Thyroid | Pass | 0.023333 | High | 23 | 1.7 - 4.3 µL | 40-100ng
Melanoma 1 | Pass | 0.01803 | High | 27 | 1.5 - 3.7 µL | 40-100ng
Paired Breast Tumor | Pass | 0.050442 | Low | 13 |  |  | We do not recommend proceeding with this sample.
Paired Normal Breast | Pass | 0.042733 | Low | 9 | 13.3 - 27.8 µL | 120-250ng
Melanoma 2 | Pass | 0.017387 | High | 31 | 1.3 - 3.2 µL | 40-100ng
NTS (all-prep) | Pass | 0.011843 | High | 22 | 1.8 - 4.5 µL | 40-100ng
N30 (all-prep) | Pass | 0.0069867 | High | 52 | 0.8 - 1.9 µL | 40-100ng
PLT | Pass | 0.02332 | High | 14 | 2.9 - 7.1 µL | 40-100ng
PLN | Pass | 0.02377 | High | 5 | 8.0 - 20.0 µL | 40-100ng
FFPE Melanoma | Pass | 0.03355 | High | 11 | 3.6 - 9.1 µL | 40-100ng
FFPE Colon | Pass | 0.019537 | High | 50 | 0.8 - 2.0 µL | 40-100ng
N19 | Pass | 0.019033 | High | 14 | 2.9 - 7.1 µL | 40-100ng
N10 | Pass | 0.018553 | High | 20 | 2.0 - 5.0 µL | 40-100ng
NA12878 | Pass | -0.0011833 | High | 20 | 2.0 - 5.0 µL | 40-100ng
Thyroid Tumor | Pass | 0.06031 | Low | 36 | 3.3 - 6.9 µL | 120-250ng
Normal Thyroid | Pass | 0.023333 | High | 23 | 1.7 - 4.3 µL | 40-100ng
Melanoma 1 | Pass | 0.01803 | High | 27 | 1.5 - 3.7 µL | 40-100ng
Paired Breast Tumor | Pass | 0.050442 | Low | 13 |  |  | We do not recommend proceeding with this sample.
Paired Normal Breast | Pass | 0.042733 | Low | 9 | 13.3 - 27.8 µL | 120-250ng
Melanoma 2 | Pass | 0.017387 | High | 31 | 1.3 - 3.2 µL | 40-100ng
NTS (all-prep) | Pass | 0.011843 | High | 22 | 1.8 - 4.5 µL | 40-100ng
N30 (all-prep) | Pass | 0.0069867 | High | 52 | 0.8 - 1.9 µL | 40-100ng
PLT | Pass | 0.02332 | High | 14 | 2.9 - 7.1 µL | 40-100ng
PLN | Pass | 0.02377 | High | 5 | 8.0 - 20.0 µL | 40-100ng
FFPE Melanoma | Pass | 0.03355 | High | 11 | 3.6 - 9.1 µL | 40-100ng
FFPE Colon | Pass | 0.019537 | High | 50 | 0.8 - 2.0 µL | 40-100ng
N19 | Pass | 0.019033 | High | 14 | 2.9 - 7.1 µL | 40-100ng
N10 | Pass | 0.018553 | High | 20 | 2.0 - 5.0 µL | 40-100ng
NA12878 | Pass | -0.0011833 | High | 20 | 2.0 - 5.0 µL | 40-100ng
Thyroid Tumor | Pass | 0.06031 | Low | 36 | 3.3 - 6.9 µL | 120-250ng
Normal Thyroid | Pass | 0.023333 | High | 23 | 1.7 - 4.3 µL | 40-100ng
Melanoma 1 | Pass | 0.01803 | High | 27 | 1.5 - 3.7 µL | 40-100ng
Paired Breast Tumor | Pass | 0.050442 | Low | 13 |  |  | We do not recommend proceeding with this sample.
Paired Normal Breast | Pass | 0.042733 | Low | 9 | 13.3 - 27.8 µL | 120-250ng
Melanoma 2 | Pass | 0.017387 | High | 31 | 1.3 - 3.2 µL | 40-100ng
NTS (all-prep) | Pass | 0.011843 | High | 22 | 1.8 - 4.5 µL | 40-100ng
N30 (all-prep) | Pass | 0.0069867 | High | 52 | 0.8 - 1.9 µL | 40-100ng
PLT | Pass | 0.02332 | High | 14 | 2.9 - 7.1 µL | 40-100ng
PLN | Pass | 0.02377 | High | 5 | 8.0 - 20.0 µL | 40-100ng
FFPE Melanoma | Pass | 0.03355 | High | 11 | 3.6 - 9.1 µL | 40-100ng
FFPE Colon | Pass | 0.019537 | High | 50 | 0.8 - 2.0 µL | 40-100ng
N19 | Pass | 0.019033 | High | 14 | 2.9 - 7.1 µL | 40-100ng
N10 | Pass | 0.018553 | High | 20 | 2.0 - 5.0 µL | 40-100ng
NA12878 | Pass | -0.0011833 | High | 20 | 2.0 - 5.0 µL | 40-100ng
Thyroid Tumor | Pass | 0.06031 | Low | 36 | 3.3 - 6.9 µL | 120-250ng
Normal Thyroid | Pass | 0.023333 | High | 23 | 1.7 - 4.3 µL | 40-100ng
Melanoma 1 | Pass | 0.01803 | High | 27 | 1.5 - 3.7 µL | 40-100ng
Paired Breast Tumor | Pass | 0.050442 | Low | 13 |  |  | We do not recommend proceeding with this sample.
Paired Normal Breast | Pass | 0.042733 | Low | 9 | 13.3 - 27.8 µL | 120-250ng
Melanoma 2 | Pass | 0.017387 | High | 31 | 1.3 - 3.2 µL | 40-100ng
NTS (all-prep) | Pass | 0.011843 | High | 22 | 1.8 - 4.5 µL | 40-100ng
N30 (all-prep) | Pass | 0.0069867 | High | 52 | 0.8 - 1.9 µL | 40-100ng
PLT | Pass | 0.02332 | High | 14 | 2.9 - 7.1 µL | 40-100ng
PLN | Pass | 0.02377 | High | 5 | 8.0 - 20.0 µL | 40-100ng
FFPE Melanoma | Pass | 0.03355 | High | 11 | 3.6 - 9.1 µL | 40-100ng
FFPE Colon | Pass | 0.019537 | High | 50 | 0.8 - 2.0 µL | 40-100ng
N19 | Pass | 0.019033 | High | 14 | 2.9 - 7.1 µL | 40-100ng
N10 | Pass | 0.018553 | High | 20 | 2.0 - 5.0 µL | 40-100ng
NA12878 | Pass | -0.0011833 | High | 20 | 2.0 - 5.0 µL | 40-100ng
Thyroid Tumor | Pass | 0.06031 | Low | 36 | 3.3 - 6.9 µL | 120-250ng
Normal Thyroid | Pass | 0.023333 | High | 23 | 1.7 - 4.3 µL | 40-100ng
Melanoma 1 | Pass | 0.01803 | High | 27 | 1.5 - 3.7 µL | 40-100ng
Control | Pass

## Calculation
Well | Raw Data | Layout | Control Position | Assays | Assay Group (bp) | Sample Position | Average CT | Standard Deviation | Control CT | ΔCT (Sample-Control) |  | 100 | 200 | Slope | X | Y
A01 | 23.328 | S1 | P10 | Assay100 | 100 | S1_100 | 23.351 | 0.039837 | 18.6033 | 4.7477 | Sample | ∆CT100 | ∆CT200 |  | ΔCT150 | ng/ml-amplifiable | Added Cycles | Lower Vol. | Upper Vol. | Sample | 100 | 200 | ∆CT | Assay QC
A02 | 23.397 | S1 | P11 | Assay100 | [6 empty] | S1 | 4.7477 | 10.7787 | 0.06031 | 7.7632 | 0.023016 | 6 | 3.3333 | 6.9444 | S1 | 23.351 | 29.9747 | 6.6237 | Pass
A03 | 23.328 | S1 | P12 | Assay100 | [6 empty] | S2 | 0.716 | 3.0493 | 0.023333 | 1.8827 | 1.3559 | 6 | 1.7391 | 4.3478 | S2 | 19.3193 | 22.2453 | 2.926 | Pass
A04 | 19.319 | S17 | P10 | Assay100 | 100 | S17_100 | 19.3193 | 0.012504 | 18.6033 | 0.716 | S3 | -0.414 | 1.389 | 0.01803 | 0.4875 | 3.5663 | 6 | 1.4815 | 3.7037 | S3 | 18.1893 | 20.585 | 2.3957 | Pass
A05 | 19.332 | S17 | P11 | Assay100 | [6 empty] | S4 | 7.2453 | 12.2895 | 0.050442 | 9.7674 | 0.005737 |  |  |  | S4 | 25.8487 | 31.4855 | 5.6368 | Pass
A06 | 19.307 | S17 | P12 | Assay100 | [6 empty] | S5 | 6.2803 | 10.5537 | 0.042733 | 8.417 | 0.014629 | 6 | 13.3333 | 27.7778 | S5 | 24.8837 | 29.7497 | 4.866 | Pass
A07 | 18.198 | S33 | P10 | Assay100 | 100 | S33_100 | 18.1893 | 0.029956 | 18.6033 | -0.414 | S6 | -0.554 | 1.1847 | 0.017387 | 0.31533 | 4.0183 | 6 | 1.2903 | 3.2258 | S6 | 18.0493 | 20.3807 | 2.3313 | Pass
A08 | 18.156 | S33 | P11 | Assay100 | [6 empty] | S7 | -0.825 | 0.35933 | 0.011843 | -0.23283 | 5.8757 | 6 | 1.8182 | 4.5455 | S7 | 17.7783 | 19.5553 | 1.777 | Pass
A09 | 18.214 | S33 | P12 | Assay100 | [6 empty] | S8 | -2.2547 | -1.556 | 0.0069867 | -1.9053 | 18.7298 | 6 | 0.76923 | 1.9231 | S8 | 16.3487 | 17.64 | 1.2913 | Pass
A10 | 25.905 | S49 | P10 | Assay100 | 100 | S49_100 | 25.8487 | 0.058621 | 18.6033 | 7.2453 | S9 | 1.041 | 3.373 | 0.02332 | 2.207 | 1.0829 | 6 | 2.8571 | 7.1429 | S9 | 19.6443 | 22.569 | 2.9247 | Pass
A11 | 25.788 | S49 | P11 | Assay100 | [6 empty] | S10 | 2.4287 | 4.8057 | 0.02377 | 3.6172 | 0.40747 | 6 | 8 | 20 | S10 | 21.032 | 24.0017 | 2.9697 | Pass
A12 | 25.853 | S49 | P12 | Assay100 | [6 empty] | S11 | 2.6727 | 6.0277 | 0.03355 | 4.3502 | 0.24515 | 6 | 3.6364 | 9.0909 | S11 | 21.276 | 25.2237 | 3.9477 | Pass
A13 | 30.212 | S1 | P22 | Assay200 | 200 | S1_200 | 29.9747 | 0.20566 | 19.196 | 10.7787 | S12 | -1.0637 | 0.89 | 0.019537 | -0.086833 | 5.3102 | 6 | 0.8 | 2 | S12 | 17.5397 | 20.086 | 2.5463 | Pass
A14 | 29.863 | S1 | P23 | Assay200 | [6 empty] | S13 | 0.254 | 2.1573 | 0.019033 | 1.2057 | 2.1678 | 6 | 2.8571 | 7.1429 | S13 | 18.8573 | 21.3533 | 2.496 | Pass
A15 | 29.849 | S1 | P24 | Assay200 | [6 empty] | S14 | 0.24167 | 2.097 | 0.018553 | 1.1693 | 2.2231 | 6 | 2 | 5 | S14 | 18.845 | 21.293 | 2.448 | Pass
A16 | 22.251 | S17 | P22 | Assay200 | 200 | S17_200 | 22.2453 | 0.0055074 | 19.196 | 3.0493 | S15 | -1.935 | -2.0533 | -0.0011833 | -1.9942 | 19.9193 | 0 | 2 | 5 | S15 | 16.6683 | 17.1427 | 0.47433 | Pass
A17 | 22.24 | S17 | P23 | Assay200 | [6 empty] | S16 | 4.7477 | 10.7787 | 0.06031 | 7.7632 | 0.023016 | 6 | 3.3333 | 6.9444 | S16 | 23.351 | 29.9747 | 6.6237 | Pass
A18 | 22.245 | S17 | P24 | Assay200 | [6 empty] | S17 | 0.716 | 3.0493 | 0.023333 | 1.8827 | 1.3559 | 6 | 1.7391 | 4.3478 | S17 | 19.3193 | 22.2453 | 2.926 | Pass
A19 | 20.559 | S33 | P22 | Assay200 | 200 | S33_200 | 20.585 | 0.032187 | 19.196 | 1.389 | S18 | -0.414 | 1.389 | 0.01803 | 0.4875 | 3.5663 | 6 | 1.4815 | 3.7037 | S18 | 18.1893 | 20.585 | 2.3957 | Pass
A20 | 20.575 | S33 | P23 | Assay200 | [6 empty] | S19 | 7.2453 | 12.2895 | 0.050442 | 9.7674 | 0.005737 |  |  |  | S19 | 25.8487 | 31.4855 | 5.6368 | Pass
A21 | 20.621 | S33 | P24 | Assay200 | [6 empty] | S20 | 6.2803 | 10.5537 | 0.042733 | 8.417 | 0.014629 | 6 | 13.3333 | 27.7778 | S20 | 24.8837 | 29.7497 | 4.866 | Pass
A22 |  | S49 | P22 | Assay200 | 200 | S49_200 | 31.4855 | 0.16758 | 19.196 | 12.2895 | S21 | -0.554 | 1.1847 | 0.017387 | 0.31533 | 4.0183 | 6 | 1.2903 | 3.2258 | S21 | 18.0493 | 20.3807 | 2.3313 | Pass
A23 | 31.604 | S49 | P23 | Assay200 | [6 empty] | S22 | -0.825 | 0.35933 | 0.011843 | -0.23283 | 5.8757 | 6 | 1.8182 | 4.5455 | S22 | 17.7783 | 19.5553 | 1.777 | Pass
A24 | 31.367 | S49 | P24 | Assay200 | [6 empty] | S23 | -2.2547 | -1.556 | 0.0069867 | -1.9053 | 18.7298 | 6 | 0.76923 | 1.9231 | S23 | 16.3487 | 17.64 | 1.2913 | Pass
B01 | 19.319 | S2 | P10 | Assay100 | 100 | S2_100 | 19.3193 | 0.012504 | 18.6033 | 0.716 | S24 | 1.041 | 3.373 | 0.02332 | 2.207 | 1.0829 | 6 | 2.8571 | 7.1429 | S24 | 19.6443 | 22.569 | 2.9247 | Pass
B02 | 19.332 | S2 | P11 | Assay100 | [6 empty] | S25 | 2.4287 | 4.8057 | 0.02377 | 3.6172 | 0.40747 | 6 | 8 | 20 | S25 | 21.032 | 24.0017 | 2.9697 | Pass
B03 | 19.307 | S2 | P12 | Assay100 | [6 empty] | S26 | 2.6727 | 6.0277 | 0.03355 | 4.3502 | 0.24515 | 6 | 3.6364 | 9.0909 | S26 | 21.276 | 25.2237 | 3.9477 | Pass
B04 | 18.198 | S18 | P10 | Assay100 | 100 | S18_100 | 18.1893 | 0.029956 | 18.6033 | -0.414 | S27 | -1.0637 | 0.89 | 0.019537 | -0.086833 | 5.3102 | 6 | 0.8 | 2 | S27 | 17.5397 | 20.086 | 2.5463 | Pass
B05 | 18.156 | S18 | P11 | Assay100 | [6 empty] | S28 | 0.254 | 2.1573 | 0.019033 | 1.2057 | 2.1678 | 6 | 2.8571 | 7.1429 | S28 | 18.8573 | 21.3533 | 2.496 | Pass
B06 | 18.214 | S18 | P12 | Assay100 | [6 empty] | S29 | 0.24167 | 2.097 | 0.018553 | 1.1693 | 2.2231 | 6 | 2 | 5 | S29 | 18.845 | 21.293 | 2.448 | Pass
B07 | 25.905 | S34 | P10 | Assay100 | 100 | S34_100 | 25.8487 | 0.058621 | 18.6033 | 7.2453 | S30 | -1.935 | -2.0533 | -0.0011833 | -1.9942 | 19.9193 | 0 | 2 | 5 | S30 | 16.6683 | 17.1427 | 0.47433 | Pass
B08 | 25.788 | S34 | P11 | Assay100 | [6 empty] | S31 | 4.7477 | 10.7787 | 0.06031 | 7.7632 | 0.023016 | 6 | 3.3333 | 6.9444 | S31 | 23.351 | 29.9747 | 6.6237 | Pass
B09 | 25.853 | S34 | P12 | Assay100 | [6 empty] | S32 | 0.716 | 3.0493 | 0.023333 | 1.8827 | 1.3559 | 6 | 1.7391 | 4.3478 | S32 | 19.3193 | 22.2453 | 2.926 | Pass
B10 | 24.839 | S50 | P10 | Assay100 | 100 | S50_100 | 24.8837 | 0.040808 | 18.6033 | 6.2803 | S33 | -0.414 | 1.389 | 0.01803 | 0.4875 | 3.5663 | 6 | 1.4815 | 3.7037 | S33 | 18.1893 | 20.585 | 2.3957 | Pass
B11 | 24.919 | S50 | P11 | Assay100 | [6 empty] | S34 | 7.2453 | 12.2895 | 0.050442 | 9.7674 | 0.005737 |  |  |  | S34 | 25.8487 | 31.4855 | 5.6368 | Pass
B12 | 24.893 | S50 | P12 | Assay100 | [6 empty] | S35 | 6.2803 | 10.5537 | 0.042733 | 8.417 | 0.014629 | 6 | 13.3333 | 27.7778 | S35 | 24.8837 | 29.7497 | 4.866 | Pass
B13 | 22.251 | S2 | P22 | Assay200 | 200 | S2_200 | 22.2453 | 0.0055074 | 19.196 | 3.0493 | S36 | -0.554 | 1.1847 | 0.017387 | 0.31533 | 4.0183 | 6 | 1.2903 | 3.2258 | S36 | 18.0493 | 20.3807 | 2.3313 | Pass
B14 | 22.24 | S2 | P23 | Assay200 | [6 empty] | S37 | -0.825 | 0.35933 | 0.011843 | -0.23283 | 5.8757 | 6 | 1.8182 | 4.5455 | S37 | 17.7783 | 19.5553 | 1.777 | Pass
B15 | 22.245 | S2 | P24 | Assay200 | [6 empty] | S38 | -2.2547 | -1.556 | 0.0069867 | -1.9053 | 18.7298 | 6 | 0.76923 | 1.9231 | S38 | 16.3487 | 17.64 | 1.2913 | Pass
B16 | 20.559 | S18 | P22 | Assay200 | 200 | S18_200 | 20.585 | 0.032187 | 19.196 | 1.389 | S39 | 1.041 | 3.373 | 0.02332 | 2.207 | 1.0829 | 6 | 2.8571 | 7.1429 | S39 | 19.6443 | 22.569 | 2.9247 | Pass
B17 | 20.575 | S18 | P23 | Assay200 | [6 empty] | S40 | 2.4287 | 4.8057 | 0.02377 | 3.6172 | 0.40747 | 6 | 8 | 20 | S40 | 21.032 | 24.0017 | 2.9697 | Pass
B18 | 20.621 | S18 | P24 | Assay200 | [6 empty] | S41 | 2.6727 | 6.0277 | 0.03355 | 4.3502 | 0.24515 | 6 | 3.6364 | 9.0909 | S41 | 21.276 | 25.2237 | 3.9477 | Pass
B19 |  | S34 | P22 | Assay200 | 200 | S34_200 | 31.4855 | 0.16758 | 19.196 | 12.2895 | S42 | -1.0637 | 0.89 | 0.019537 | -0.086833 | 5.3102 | 6 | 0.8 | 2 | S42 | 17.5397 | 20.086 | 2.5463 | Pass
B20 | 31.604 | S34 | P23 | Assay200 | [6 empty] | S43 | 0.254 | 2.1573 | 0.019033 | 1.2057 | 2.1678 | 6 | 2.8571 | 7.1429 | S43 | 18.8573 | 21.3533 | 2.496 | Pass
B21 | 31.367 | S34 | P24 | Assay200 | [6 empty] | S44 | 0.24167 | 2.097 | 0.018553 | 1.1693 | 2.2231 | 6 | 2 | 5 | S44 | 18.845 | 21.293 | 2.448 | Pass
B22 | 29.804 | S50 | P22 | Assay200 | 200 | S50_200 | 29.7497 | 0.28146 | 19.196 | 10.5537 | S45 | -1.935 | -2.0533 | -0.0011833 | -1.9942 | 19.9193 | 0 | 2 | 5 | S45 | 16.6683 | 17.1427 | 0.47433 | Pass
B23 | 29.445 | S50 | P23 | Assay200 | [6 empty] | S46 | 4.7477 | 10.7787 | 0.06031 | 7.7632 | 0.023016 | 6 | 3.3333 | 6.9444 | S46 | 23.351 | 29.9747 | 6.6237 | Pass
B24 | 30 | S50 | P24 | Assay200 | [6 empty] | S47 | 0.716 | 3.0493 | 0.023333 | 1.8827 | 1.3559 | 6 | 1.7391 | 4.3478 | S47 | 19.3193 | 22.2453 | 2.926 | Pass
C01 | 18.198 | S3 | P10 | Assay100 | 100 | S3_100 | 18.1893 | 0.029956 | 18.6033 | -0.414 | S48 | -0.414 | 1.389 | 0.01803 | 0.4875 | 3.5663 | 6 | 1.4815 | 3.7037 | S48 | 18.1893 | 20.585 | 2.3957 | Pass
C02 | 18.156 | S3 | P11 | Assay100 | [6 empty] | S49 | 7.2453 | 12.2895 | 0.050442 | 9.7674 | 0.005737 |  |  |  | S49 | 25.8487 | 31.4855 | 5.6368 | Pass
C03 | 18.214 | S3 | P12 | Assay100 | [6 empty] | S50 | 6.2803 | 10.5537 | 0.042733 | 8.417 | 0.014629 | 6 | 13.3333 | 27.7778 | S50 | 24.8837 | 29.7497 | 4.866 | Pass
C04 | 25.905 | S19 | P10 | Assay100 | 100 | S19_100 | 25.8487 | 0.058621 | 18.6033 | 7.2453 | S51 | -0.554 | 1.1847 | 0.017387 | 0.31533 | 4.0183 | 6 | 1.2903 | 3.2258 | S51 | 18.0493 | 20.3807 | 2.3313 | Pass
C05 | 25.788 | S19 | P11 | Assay100 | [6 empty] | S52 | -0.825 | 0.35933 | 0.011843 | -0.23283 | 5.8757 | 6 | 1.8182 | 4.5455 | S52 | 17.7783 | 19.5553 | 1.777 | Pass
C06 | 25.853 | S19 | P12 | Assay100 | [6 empty] | S53 | -2.2547 | -1.556 | 0.0069867 | -1.9053 | 18.7298 | 6 | 0.76923 | 1.9231 | S53 | 16.3487 | 17.64 | 1.2913 | Pass
C07 | 24.839 | S35 | P10 | Assay100 | 100 | S35_100 | 24.8837 | 0.040808 | 18.6033 | 6.2803 | S54 | 1.041 | 3.373 | 0.02332 | 2.207 | 1.0829 | 6 | 2.8571 | 7.1429 | S54 | 19.6443 | 22.569 | 2.9247 | Pass
C08 | 24.919 | S35 | P11 | Assay100 | [6 empty] | S55 | 2.4287 | 4.8057 | 0.02377 | 3.6172 | 0.40747 | 6 | 8 | 20 | S55 | 21.032 | 24.0017 | 2.9697 | Pass
C09 | 24.893 | S35 | P12 | Assay100 | [6 empty] | S56 | 2.6727 | 6.0277 | 0.03355 | 4.3502 | 0.24515 | 6 | 3.6364 | 9.0909 | S56 | 21.276 | 25.2237 | 3.9477 | Pass
C10 | 18.062 | S51 | P10 | Assay100 | 100 | S51_100 | 18.0493 | 0.035725 | 18.6033 | -0.554 | S57 | -1.0637 | 0.89 | 0.019537 | -0.086833 | 5.3102 | 6 | 0.8 | 2 | S57 | 17.5397 | 20.086 | 2.5463 | Pass
C11 | 18.077 | S51 | P11 | Assay100 | [6 empty] | S58 | 0.254 | 2.1573 | 0.019033 | 1.2057 | 2.1678 | 6 | 2.8571 | 7.1429 | S58 | 18.8573 | 21.3533 | 2.496 | Pass
C12 | 18.009 | S51 | P12 | Assay100 | [6 empty] | S59 | 0.24167 | 2.097 | 0.018553 | 1.1693 | 2.2231 | 6 | 2 | 5 | S59 | 18.845 | 21.293 | 2.448 | Pass
C13 | 20.559 | S3 | P22 | Assay200 | 200 | S3_200 | 20.585 | 0.032187 | 19.196 | 1.389 | S60 | -1.935 | -2.0533 | -0.0011833 | -1.9942 | 19.9193 | 0 | 2 | 5 | S60 | 16.6683 | 17.1427 | 0.47433 | Pass
C14 | 20.575 | S3 | P23 | Assay200 | [6 empty] | S61 | 4.7477 | 10.7787 | 0.06031 | 7.7632 | 0.023016 | 6 | 3.3333 | 6.9444 | S61 | 23.351 | 29.9747 | 6.6237 | Pass
C15 | 20.621 | S3 | P24 | Assay200 | [6 empty] | S62 | 0.716 | 3.0493 | 0.023333 | 1.8827 | 1.3559 | 6 | 1.7391 | 4.3478 | S62 | 19.3193 | 22.2453 | 2.926 | Pass
C16 |  | S19 | P22 | Assay200 | 200 | S19_200 | 31.4855 | 0.16758 | 19.196 | 12.2895 | S63 | -0.414 | 1.389 | 0.01803 | 0.4875 | 3.5663 | 6 | 1.4815 | 3.7037 | S63 | 18.1893 | 20.585 | 2.3957 | Pass
C17 | 31.604 | S19 | P23 | Assay200 | [6 empty] | Control | 0 | 0 | 0 | 0 | 5 | 0 |  |  | Control | 18.6033 | 19.196 | 0.59267 | Pass
C18 | 31.367 | S19 | P24 | Assay200
C19 | 29.804 | S35 | P22 | Assay200 | 200 | S35_200 | 29.7497 | 0.28146 | 19.196 | 10.5537 |  |  |  |  | If QC score<=0.04, use 40 - 100ng of the sample
C20 | 29.445 | S35 | P23 | Assay200 | [10 empty] | If QC score>0.04 and X<=9, use 120 - 250ng of the sample
C21 | 30 | S35 | P24 | Assay200 | [10 empty] | If X>9, do not proceed with sequencing sample
C22 | 20.367 | S51 | P22 | Assay200 | 200 | S51_200 | 20.3807 | 0.031786 | 19.196 | 1.1847
C23 | 20.417 | S51 | P23 | Assay200
C24 | 20.358 | S51 | P24 | Assay200
D01 | 25.905 | S4 | P10 | Assay100 | 100 | S4_100 | 25.8487 | 0.058621 | 18.6033 | 7.2453
D02 | 25.788 | S4 | P11 | Assay100
D03 | 25.853 | S4 | P12 | Assay100
D04 | 24.839 | S20 | P10 | Assay100 | 100 | S20_100 | 24.8837 | 0.040808 | 18.6033 | 6.2803
D05 | 24.919 | S20 | P11 | Assay100
D06 | 24.893 | S20 | P12 | Assay100
D07 | 18.062 | S36 | P10 | Assay100 | 100 | S36_100 | 18.0493 | 0.035725 | 18.6033 | -0.554
D08 | 18.077 | S36 | P11 | Assay100
D09 | 18.009 | S36 | P12 | Assay100
D10 | 17.841 | S52 | P10 | Assay100 | 100 | S52_100 | 17.7783 | 0.055374 | 18.6033 | -0.825
D11 | 17.736 | S52 | P11 | Assay100
D12 | 17.758 | S52 | P12 | Assay100
D13 |  | S4 | P22 | Assay200 | 200 | S4_200 | 31.4855 | 0.16758 | 19.196 | 12.2895
D14 | 31.604 | S4 | P23 | Assay200
D15 | 31.367 | S4 | P24 | Assay200
D16 | 29.804 | S20 | P22 | Assay200 | 200 | S20_200 | 29.7497 | 0.28146 | 19.196 | 10.5537
D17 | 29.445 | S20 | P23 | Assay200
D18 | 30 | S20 | P24 | Assay200
D19 | 20.367 | S36 | P22 | Assay200 | 200 | S36_200 | 20.3807 | 0.031786 | 19.196 | 1.1847
D20 | 20.417 | S36 | P23 | Assay200
D21 | 20.358 | S36 | P24 | Assay200
D22 | 19.567 | S52 | P22 | Assay200 | 200 | S52_200 | 19.5553 | 0.015308 | 19.196 | 0.35933
D23 | 19.561 | S52 | P23 | Assay200
D24 | 19.538 | S52 | P24 | Assay200
E01 | 24.839 | S5 | P10 | Assay100 | 100 | S5_100 | 24.8837 | 0.040808 | 18.6033 | 6.2803
E02 | 24.919 | S5 | P11 | Assay100
E03 | 24.893 | S5 | P12 | Assay100
E04 | 18.062 | S21 | P10 | Assay100 | 100 | S21_100 | 18.0493 | 0.035725 | 18.6033 | -0.554
E05 | 18.077 | S21 | P11 | Assay100
E06 | 18.009 | S21 | P12 | Assay100
E07 | 17.841 | S37 | P10 | Assay100 | 100 | S37_100 | 17.7783 | 0.055374 | 18.6033 | -0.825
E08 | 17.736 | S37 | P11 | Assay100
E09 | 17.758 | S37 | P12 | Assay100
E10 | 16.369 | S53 | P10 | Assay100 | 100 | S53_100 | 16.3487 | 0.01877 | 18.6033 | -2.2547
E11 | 16.345 | S53 | P11 | Assay100
E12 | 16.332 | S53 | P12 | Assay100
E13 | 29.804 | S5 | P22 | Assay200 | 200 | S5_200 | 29.7497 | 0.28146 | 19.196 | 10.5537
E14 | 29.445 | S5 | P23 | Assay200
E15 | 30 | S5 | P24 | Assay200
E16 | 20.367 | S21 | P22 | Assay200 | 200 | S21_200 | 20.3807 | 0.031786 | 19.196 | 1.1847
E17 | 20.417 | S21 | P23 | Assay200
E18 | 20.358 | S21 | P24 | Assay200
E19 | 19.567 | S37 | P22 | Assay200 | 200 | S37_200 | 19.5553 | 0.015308 | 19.196 | 0.35933
E20 | 19.561 | S37 | P23 | Assay200
E21 | 19.538 | S37 | P24 | Assay200
E22 | 17.655 | S53 | P22 | Assay200 | 200 | S53_200 | 17.64 | 0.014526 | 19.196 | -1.556
E23 | 17.626 | S53 | P23 | Assay200
E24 | 17.639 | S53 | P24 | Assay200
F01 | 18.062 | S6 | P10 | Assay100 | 100 | S6_100 | 18.0493 | 0.035725 | 18.6033 | -0.554
F02 | 18.077 | S6 | P11 | Assay100
F03 | 18.009 | S6 | P12 | Assay100
F04 | 17.841 | S22 | P10 | Assay100 | 100 | S22_100 | 17.7783 | 0.055374 | 18.6033 | -0.825
F05 | 17.736 | S22 | P11 | Assay100
F06 | 17.758 | S22 | P12 | Assay100
F07 | 16.369 | S38 | P10 | Assay100 | 100 | S38_100 | 16.3487 | 0.01877 | 18.6033 | -2.2547
F08 | 16.345 | S38 | P11 | Assay100
F09 | 16.332 | S38 | P12 | Assay100
F10 | 19.644 | S54 | P10 | Assay100 | 100 | S54_100 | 19.6443 | 0.013503 | 18.6033 | 1.041
F11 | 19.631 | S54 | P11 | Assay100
F12 | 19.658 | S54 | P12 | Assay100
F13 | 20.367 | S6 | P22 | Assay200 | 200 | S6_200 | 20.3807 | 0.031786 | 19.196 | 1.1847
F14 | 20.417 | S6 | P23 | Assay200
F15 | 20.358 | S6 | P24 | Assay200
F16 | 19.567 | S22 | P22 | Assay200 | 200 | S22_200 | 19.5553 | 0.015308 | 19.196 | 0.35933
F17 | 19.561 | S22 | P23 | Assay200
F18 | 19.538 | S22 | P24 | Assay200
F19 | 17.655 | S38 | P22 | Assay200 | 200 | S38_200 | 17.64 | 0.014526 | 19.196 | -1.556
F20 | 17.626 | S38 | P23 | Assay200
F21 | 17.639 | S38 | P24 | Assay200
F22 | 22.575 | S54 | P22 | Assay200 | 200 | S54_200 | 22.569 | 0.0060005 | 19.196 | 3.373
F23 | 22.563 | S54 | P23 | Assay200
F24 | 22.569 | S54 | P24 | Assay200
G01 | 17.841 | S7 | P10 | Assay100 | 100 | S7_100 | 17.7783 | 0.055374 | 18.6033 | -0.825
G02 | 17.736 | S7 | P11 | Assay100
G03 | 17.758 | S7 | P12 | Assay100
G04 | 16.369 | S23 | P10 | Assay100 | 100 | S23_100 | 16.3487 | 0.01877 | 18.6033 | -2.2547
G05 | 16.345 | S23 | P11 | Assay100
G06 | 16.332 | S23 | P12 | Assay100
G07 | 19.644 | S39 | P10 | Assay100 | 100 | S39_100 | 19.6443 | 0.013503 | 18.6033 | 1.041
G08 | 19.631 | S39 | P11 | Assay100
G09 | 19.658 | S39 | P12 | Assay100
G10 | 21.098 | S55 | P10 | Assay100 | 100 | S55_100 | 21.032 | 0.057472 | 18.6033 | 2.4287
G11 | 20.993 | S55 | P11 | Assay100
G12 | 21.005 | S55 | P12 | Assay100
G13 | 19.567 | S7 | P22 | Assay200 | 200 | S7_200 | 19.5553 | 0.015308 | 19.196 | 0.35933
G14 | 19.561 | S7 | P23 | Assay200
G15 | 19.538 | S7 | P24 | Assay200
G16 | 17.655 | S23 | P22 | Assay200 | 200 | S23_200 | 17.64 | 0.014526 | 19.196 | -1.556
G17 | 17.626 | S23 | P23 | Assay200
G18 | 17.639 | S23 | P24 | Assay200
G19 | 22.575 | S39 | P22 | Assay200 | 200 | S39_200 | 22.569 | 0.0060005 | 19.196 | 3.373
G20 | 22.563 | S39 | P23 | Assay200
G21 | 22.569 | S39 | P24 | Assay200
G22 | 23.975 | S55 | P22 | Assay200 | 200 | S55_200 | 24.0017 | 0.074662 | 19.196 | 4.8057
G23 | 23.944 | S55 | P23 | Assay200
G24 | 24.086 | S55 | P24 | Assay200
H01 | 16.369 | S8 | P10 | Assay100 | 100 | S8_100 | 16.3487 | 0.01877 | 18.6033 | -2.2547
H02 | 16.345 | S8 | P11 | Assay100
H03 | 16.332 | S8 | P12 | Assay100
H04 | 19.644 | S24 | P10 | Assay100 | 100 | S24_100 | 19.6443 | 0.013503 | 18.6033 | 1.041
H05 | 19.631 | S24 | P11 | Assay100
H06 | 19.658 | S24 | P12 | Assay100
H07 | 21.098 | S40 | P10 | Assay100 | 100 | S40_100 | 21.032 | 0.057472 | 18.6033 | 2.4287
H08 | 20.993 | S40 | P11 | Assay100
H09 | 21.005 | S40 | P12 | Assay100
H10 | 21.302 | S56 | P10 | Assay100 | 100 | S56_100 | 21.276 | 0.027056 | 18.6033 | 2.6727
H11 | 21.278 | S56 | P11 | Assay100
H12 | 21.248 | S56 | P12 | Assay100
H13 | 17.655 | S8 | P22 | Assay200 | 200 | S8_200 | 17.64 | 0.014526 | 19.196 | -1.556
H14 | 17.626 | S8 | P23 | Assay200
H15 | 17.639 | S8 | P24 | Assay200
H16 | 22.575 | S24 | P22 | Assay200 | 200 | S24_200 | 22.569 | 0.0060005 | 19.196 | 3.373
H17 | 22.563 | S24 | P23 | Assay200
H18 | 22.569 | S24 | P24 | Assay200
H19 | 23.975 | S40 | P22 | Assay200 | 200 | S40_200 | 24.0017 | 0.074662 | 19.196 | 4.8057
H20 | 23.944 | S40 | P23 | Assay200
H21 | 24.086 | S40 | P24 | Assay200
H22 | 25.154 | S56 | P22 | Assay200 | 200 | S56_200 | 25.2237 | 0.06853 | 19.196 | 6.0277
H23 | 25.291 | S56 | P23 | Assay200
H24 | 25.226 | S56 | P24 | Assay200
I01 | 19.644 | S9 | P10 | Assay100 | 100 | S9_100 | 19.6443 | 0.013503 | 18.6033 | 1.041
I02 | 19.631 | S9 | P11 | Assay100
I03 | 19.658 | S9 | P12 | Assay100
I04 | 21.098 | S25 | P10 | Assay100 | 100 | S25_100 | 21.032 | 0.057472 | 18.6033 | 2.4287
I05 | 20.993 | S25 | P11 | Assay100
I06 | 21.005 | S25 | P12 | Assay100
I07 | 21.302 | S41 | P10 | Assay100 | 100 | S41_100 | 21.276 | 0.027056 | 18.6033 | 2.6727
I08 | 21.278 | S41 | P11 | Assay100
I09 | 21.248 | S41 | P12 | Assay100
I10 | 17.584 | S57 | P10 | Assay100 | 100 | S57_100 | 17.5397 | 0.039323 | 18.6033 | -1.0637
I11 | 17.526 | S57 | P11 | Assay100
I12 | 17.509 | S57 | P12 | Assay100
I13 | 22.575 | S9 | P22 | Assay200 | 200 | S9_200 | 22.569 | 0.0060005 | 19.196 | 3.373
I14 | 22.563 | S9 | P23 | Assay200
I15 | 22.569 | S9 | P24 | Assay200
I16 | 23.975 | S25 | P22 | Assay200 | 200 | S25_200 | 24.0017 | 0.074662 | 19.196 | 4.8057
I17 | 23.944 | S25 | P23 | Assay200
I18 | 24.086 | S25 | P24 | Assay200
I19 | 25.154 | S41 | P22 | Assay200 | 200 | S41_200 | 25.2237 | 0.06853 | 19.196 | 6.0277
I20 | 25.291 | S41 | P23 | Assay200
I21 | 25.226 | S41 | P24 | Assay200
I22 | 20.004 | S57 | P22 | Assay200 | 200 | S57_200 | 20.086 | 0.071359 | 19.196 | 0.89
I23 | 20.134 | S57 | P23 | Assay200
I24 | 20.12 | S57 | P24 | Assay200
J01 | 21.098 | S10 | P10 | Assay100 | 100 | S10_100 | 21.032 | 0.057472 | 18.6033 | 2.4287
J02 | 20.993 | S10 | P11 | Assay100
J03 | 21.005 | S10 | P12 | Assay100
J04 | 21.302 | S26 | P10 | Assay100 | 100 | S26_100 | 21.276 | 0.027056 | 18.6033 | 2.6727
J05 | 21.278 | S26 | P11 | Assay100
J06 | 21.248 | S26 | P12 | Assay100
J07 | 17.584 | S42 | P10 | Assay100 | 100 | S42_100 | 17.5397 | 0.039323 | 18.6033 | -1.0637
J08 | 17.526 | S42 | P11 | Assay100
J09 | 17.509 | S42 | P12 | Assay100
J10 | 18.892 | S58 | P10 | Assay100 | 100 | S58_100 | 18.8573 | 0.043155 | 18.6033 | 0.254
J11 | 18.871 | S58 | P11 | Assay100
J12 | 18.809 | S58 | P12 | Assay100
J13 | 23.975 | S10 | P22 | Assay200 | 200 | S10_200 | 24.0017 | 0.074662 | 19.196 | 4.8057
J14 | 23.944 | S10 | P23 | Assay200
J15 | 24.086 | S10 | P24 | Assay200
J16 | 25.154 | S26 | P22 | Assay200 | 200 | S26_200 | 25.2237 | 0.06853 | 19.196 | 6.0277
J17 | 25.291 | S26 | P23 | Assay200
J18 | 25.226 | S26 | P24 | Assay200
J19 | 20.004 | S42 | P22 | Assay200 | 200 | S42_200 | 20.086 | 0.071359 | 19.196 | 0.89
J20 | 20.134 | S42 | P23 | Assay200
J21 | 20.12 | S42 | P24 | Assay200
J22 | 21.336 | S58 | P22 | Assay200 | 200 | S58_200 | 21.3533 | 0.025813 | 19.196 | 2.1573
J23 | 21.383 | S58 | P23 | Assay200
J24 | 21.341 | S58 | P24 | Assay200
K01 | 21.302 | S11 | P10 | Assay100 | 100 | S11_100 | 21.276 | 0.027056 | 18.6033 | 2.6727
K02 | 21.278 | S11 | P11 | Assay100
K03 | 21.248 | S11 | P12 | Assay100
K04 | 17.584 | S27 | P10 | Assay100 | 100 | S27_100 | 17.5397 | 0.039323 | 18.6033 | -1.0637
K05 | 17.526 | S27 | P11 | Assay100
K06 | 17.509 | S27 | P12 | Assay100
K07 | 18.892 | S43 | P10 | Assay100 | 100 | S43_100 | 18.8573 | 0.043155 | 18.6033 | 0.254
K08 | 18.871 | S43 | P11 | Assay100
K09 | 18.809 | S43 | P12 | Assay100
K10 | 18.864 | S59 | P10 | Assay100 | 100 | S59_100 | 18.845 | 0.016822 | 18.6033 | 0.24167
K11 | 18.832 | S59 | P11 | Assay100
K12 | 18.839 | S59 | P12 | Assay100
K13 | 25.154 | S11 | P22 | Assay200 | 200 | S11_200 | 25.2237 | 0.06853 | 19.196 | 6.0277
K14 | 25.291 | S11 | P23 | Assay200
K15 | 25.226 | S11 | P24 | Assay200
K16 | 20.004 | S27 | P22 | Assay200 | 200 | S27_200 | 20.086 | 0.071359 | 19.196 | 0.89
K17 | 20.134 | S27 | P23 | Assay200
K18 | 20.12 | S27 | P24 | Assay200
K19 | 21.336 | S43 | P22 | Assay200 | 200 | S43_200 | 21.3533 | 0.025813 | 19.196 | 2.1573
K20 | 21.383 | S43 | P23 | Assay200
K21 | 21.341 | S43 | P24 | Assay200
K22 | 21.26 | S59 | P22 | Assay200 | 200 | S59_200 | 21.293 | 0.029816 | 19.196 | 2.097
K23 | 21.301 | S59 | P23 | Assay200
K24 | 21.318 | S59 | P24 | Assay200
L01 | 17.584 | S12 | P10 | Assay100 | 100 | S12_100 | 17.5397 | 0.039323 | 18.6033 | -1.0637
L02 | 17.526 | S12 | P11 | Assay100
L03 | 17.509 | S12 | P12 | Assay100
L04 | 18.892 | S28 | P10 | Assay100 | 100 | S28_100 | 18.8573 | 0.043155 | 18.6033 | 0.254
L05 | 18.871 | S28 | P11 | Assay100
L06 | 18.809 | S28 | P12 | Assay100
L07 | 18.864 | S44 | P10 | Assay100 | 100 | S44_100 | 18.845 | 0.016822 | 18.6033 | 0.24167
L08 | 18.832 | S44 | P11 | Assay100
L09 | 18.839 | S44 | P12 | Assay100
L10 | 16.68 | S60 | P10 | Assay100 | 100 | S60_100 | 16.6683 | 0.012583 | 18.6033 | -1.935
L11 | 16.67 | S60 | P11 | Assay100
L12 | 16.655 | S60 | P12 | Assay100
L13 | 20.004 | S12 | P22 | Assay200 | 200 | S12_200 | 20.086 | 0.071359 | 19.196 | 0.89
L14 | 20.134 | S12 | P23 | Assay200
L15 | 20.12 | S12 | P24 | Assay200
L16 | 21.336 | S28 | P22 | Assay200 | 200 | S28_200 | 21.3533 | 0.025813 | 19.196 | 2.1573
L17 | 21.383 | S28 | P23 | Assay200
L18 | 21.341 | S28 | P24 | Assay200
L19 | 21.26 | S44 | P22 | Assay200 | 200 | S44_200 | 21.293 | 0.029816 | 19.196 | 2.097
L20 | 21.301 | S44 | P23 | Assay200
L21 | 21.318 | S44 | P24 | Assay200
L22 | 17.135 | S60 | P22 | Assay200 | 200 | S60_200 | 17.1427 | 0.018717 | 19.196 | -2.0533
L23 | 17.129 | S60 | P23 | Assay200
L24 | 17.164 | S60 | P24 | Assay200
M01 | 18.892 | S13 | P10 | Assay100 | 100 | S13_100 | 18.8573 | 0.043155 | 18.6033 | 0.254
M02 | 18.871 | S13 | P11 | Assay100
M03 | 18.809 | S13 | P12 | Assay100
M04 | 18.864 | S29 | P10 | Assay100 | 100 | S29_100 | 18.845 | 0.016822 | 18.6033 | 0.24167
M05 | 18.832 | S29 | P11 | Assay100
M06 | 18.839 | S29 | P12 | Assay100
M07 | 16.68 | S45 | P10 | Assay100 | 100 | S45_100 | 16.6683 | 0.012583 | 18.6033 | -1.935
M08 | 16.67 | S45 | P11 | Assay100
M09 | 16.655 | S45 | P12 | Assay100
M10 | 23.328 | S61 | P10 | Assay100 | 100 | S61_100 | 23.351 | 0.039837 | 18.6033 | 4.7477
M11 | 23.397 | S61 | P11 | Assay100
M12 | 23.328 | S61 | P12 | Assay100
M13 | 21.336 | S13 | P22 | Assay200 | 200 | S13_200 | 21.3533 | 0.025813 | 19.196 | 2.1573
M14 | 21.383 | S13 | P23 | Assay200
M15 | 21.341 | S13 | P24 | Assay200
M16 | 21.26 | S29 | P22 | Assay200 | 200 | S29_200 | 21.293 | 0.029816 | 19.196 | 2.097
M17 | 21.301 | S29 | P23 | Assay200
M18 | 21.318 | S29 | P24 | Assay200
M19 | 17.135 | S45 | P22 | Assay200 | 200 | S45_200 | 17.1427 | 0.018717 | 19.196 | -2.0533
M20 | 17.129 | S45 | P23 | Assay200
M21 | 17.164 | S45 | P24 | Assay200
M22 | 30.212 | S61 | P22 | Assay200 | 200 | S61_200 | 29.9747 | 0.20566 | 19.196 | 10.7787
M23 | 29.863 | S61 | P23 | Assay200
M24 | 29.849 | S61 | P24 | Assay200
N01 | 18.864 | S14 | P10 | Assay100 | 100 | S14_100 | 18.845 | 0.016822 | 18.6033 | 0.24167
N02 | 18.832 | S14 | P11 | Assay100
N03 | 18.839 | S14 | P12 | Assay100
N04 | 16.68 | S30 | P10 | Assay100 | 100 | S30_100 | 16.6683 | 0.012583 | 18.6033 | -1.935
N05 | 16.67 | S30 | P11 | Assay100
N06 | 16.655 | S30 | P12 | Assay100
N07 | 23.328 | S46 | P10 | Assay100 | 100 | S46_100 | 23.351 | 0.039837 | 18.6033 | 4.7477
N08 | 23.397 | S46 | P11 | Assay100
N09 | 23.328 | S46 | P12 | Assay100
N10 | 19.319 | S62 | P10 | Assay100 | 100 | S62_100 | 19.3193 | 0.012504 | 18.6033 | 0.716
N11 | 19.332 | S62 | P11 | Assay100
N12 | 19.307 | S62 | P12 | Assay100
N13 | 21.26 | S14 | P22 | Assay200 | 200 | S14_200 | 21.293 | 0.029816 | 19.196 | 2.097
N14 | 21.301 | S14 | P23 | Assay200
N15 | 21.318 | S14 | P24 | Assay200
N16 | 17.135 | S30 | P22 | Assay200 | 200 | S30_200 | 17.1427 | 0.018717 | 19.196 | -2.0533
N17 | 17.129 | S30 | P23 | Assay200
N18 | 17.164 | S30 | P24 | Assay200
N19 | 30.212 | S46 | P22 | Assay200 | 200 | S46_200 | 29.9747 | 0.20566 | 19.196 | 10.7787
N20 | 29.863 | S46 | P23 | Assay200
N21 | 29.849 | S46 | P24 | Assay200
N22 | 22.251 | S62 | P22 | Assay200 | 200 | S62_200 | 22.2453 | 0.0055074 | 19.196 | 3.0493
N23 | 22.24 | S62 | P23 | Assay200
N24 | 22.245 | S62 | P24 | Assay200
O01 | 16.68 | S15 | P10 | Assay100 | 100 | S15_100 | 16.6683 | 0.012583 | 18.6033 | -1.935
O02 | 16.67 | S15 | P11 | Assay100
O03 | 16.655 | S15 | P12 | Assay100
O04 | 23.328 | S31 | P10 | Assay100 | 100 | S31_100 | 23.351 | 0.039837 | 18.6033 | 4.7477
O05 | 23.397 | S31 | P11 | Assay100
O06 | 23.328 | S31 | P12 | Assay100
O07 | 19.319 | S47 | P10 | Assay100 | 100 | S47_100 | 19.3193 | 0.012504 | 18.6033 | 0.716
O08 | 19.332 | S47 | P11 | Assay100
O09 | 19.307 | S47 | P12 | Assay100
O10 | 18.198 | S63 | P10 | Assay100 | 100 | S63_100 | 18.1893 | 0.029956 | 18.6033 | -0.414
O11 | 18.156 | S63 | P11 | Assay100
O12 | 18.214 | S63 | P12 | Assay100
O13 | 17.135 | S15 | P22 | Assay200 | 200 | S15_200 | 17.1427 | 0.018717 | 19.196 | -2.0533
O14 | 17.129 | S15 | P23 | Assay200
O15 | 17.164 | S15 | P24 | Assay200
O16 | 30.212 | S31 | P22 | Assay200 | 200 | S31_200 | 29.9747 | 0.20566 | 19.196 | 10.7787
O17 | 29.863 | S31 | P23 | Assay200
O18 | 29.849 | S31 | P24 | Assay200
O19 | 22.251 | S47 | P22 | Assay200 | 200 | S47_200 | 22.2453 | 0.0055074 | 19.196 | 3.0493
O20 | 22.24 | S47 | P23 | Assay200
O21 | 22.245 | S47 | P24 | Assay200
O22 | 20.559 | S63 | P22 | Assay200 | 200 | S63_200 | 20.585 | 0.032187 | 19.196 | 1.389
O23 | 20.575 | S63 | P23 | Assay200
O24 | 20.621 | S63 | P24 | Assay200
P01 | 23.328 | S16 | P10 | Assay100 | 100 | S16_100 | 23.351 | 0.039837 | 18.6033 | 4.7477
P02 | 23.397 | S16 | P11 | Assay100
P03 | 23.328 | S16 | P12 | Assay100
P04 | 19.319 | S32 | P10 | Assay100 | 100 | S32_100 | 19.3193 | 0.012504 | 18.6033 | 0.716
P05 | 19.332 | S32 | P11 | Assay100
P06 | 19.307 | S32 | P12 | Assay100
P07 | 18.198 | S48 | P10 | Assay100 | 100 | S48_100 | 18.1893 | 0.029956 | 18.6033 | -0.414
P08 | 18.156 | S48 | P11 | Assay100
P09 | 18.214 | S48 | P12 | Assay100
P10 | 18.626 | Control | P10 | Assay100 | 100 | Control_100 | 18.6033 | 0.019857 | 18.6033 | 0
P11 | 18.589 | Control | P11 | Assay100
P12 | 18.595 | Control | P12 | Assay100
P13 | 30.212 | S16 | P22 | Assay200 | 200 | S16_200 | 29.9747 | 0.20566 | 19.196 | 10.7787
P14 | 29.863 | S16 | P23 | Assay200
P15 | 29.849 | S16 | P24 | Assay200
P16 | 22.251 | S32 | P22 | Assay200 | 200 | S32_200 | 22.2453 | 0.0055074 | 19.196 | 3.0493
P17 | 22.24 | S32 | P23 | Assay200
P18 | 22.245 | S32 | P24 | Assay200
P19 | 20.559 | S48 | P22 | Assay200 | 200 | S48_200 | 20.585 | 0.032187 | 19.196 | 1.389
P20 | 20.575 | S48 | P23 | Assay200
P21 | 20.621 | S48 | P24 | Assay200
P22 | 19.169 | Control | P22 | Assay200 | 200 | Control_200 | 19.196 | 0.02488 | 19.196 | 0
P23 | 19.201 | Control | P23 | Assay200
P24 | 19.218 | Control | P24 | Assay200