## Instructions
Instructions for analyzing GeneRead DNAseq Library Quant Array (Ion Torrent) results with this file:
Generally, only change data in yellow or orange cells. Gray and white cells contain formulae for calculation or results. Please do not change them.
1. Raw data worksheet: Copy and paste the CT values from the real-time PCR instrument runs of your experimental samples to the yellow section of the "Raw data" worksheet starting at cell B2.
2. Sample library dilution: Input your sample library dilution fold (in orange) if you use different dilution fold other than the ones recommended in the handbook.
3. Library information: Change the total primer pairs (in orange) according to corresponding GeneRead NGS DNAseq Gene Panel (Find the primer pair information in Table 1 in the same worksheet. This is critical for target enrichment QC). Input your library insert size (in orange). This will be used for size adjustment to account for the difference in size between the average amplicon length of the library and the DNA standard (183 bp).
4. QC report: All NTC should be more than 29. Slope should be within –3.1 to –3.6. Standard deviation for triplicates should be <0.4.
5. Results: Template dilution factors (dilution fold from original library) are displayed, which you can use to dilute your original libraries (to 4 pM). Use 25 µl of diluted library for emPCR, which is recommended from Ion Torrent template kit. Target enrichment procedure is monitored by using spikein primers to measure control templates. The control template amount should be in a reasonable range in the library. All libraries should display "Pass" or "Marginal" for processing to sequencing.
6. Calculations: To view the formulas and equations used to calculate the QC data and results, you may review but not alter the "Calculations" worksheet.
.

## Raw data
Well | Raw data
A01 | 9.719
A02 | 9.7991
A03 | 9.7376
A04 | 13.3176
A05 | 13.2803
A06 | 13.3032
A07 | 16.7364
A08 | 16.7909
A09 | 16.8021
A10 | 20.333
A11 | 20.4496
A12 | 20.3468
B01 | 23.8822
B02 | 23.8979
B03 | 23.8454
B04 | Undetermined
B05 | Undetermined
B06 | Undetermined
B07 | Undetermined
B08 | Undetermined
B09 | Undetermined
B10 | Undetermined
B11 | Undetermined
B12 | Undetermined
C01 | 27.6263
C02 | 28.992
C03 | 25.3635
C04 | 25.5388
C05 | 28.8413
C06 | 29.2539
C07 | 11.9539
C08 | 11.967
C09 | 11.9829
C10 | 15.4695
C11 | 15.4925
C12 | 15.5201
D01 | 21.9818
D02 | 23.2177
D03 | 22.0128
D04 | 20.4558
D05 | 22.4246
D06 | 23.3641
D07 | 8.6829
D08 | 8.6929
D09 | 8.7377
D10 | 13.3192
D11 | 13.3044
D12 | 13.2823
E01 | 11.8106
E02 | 12.4618
E03 | 11.7785
E04 | 11.7497
E05 | 12.7409
E06 | 13.156
E07 | 4.8088
E08 | 4.8192
E09 | 4.8089
E10 | 8.1884
E11 | 8.2265
E12 | 8.1931
F01 | 16.8004
F02 | 18.0028
F03 | 16.9313
F04 | 17.1467
F05 | 17.3042
F06 | 18.5281
F07 | 9.3232
F08 | 9.3416
F09 | 9.3502
F10 | 12.9285
F11 | 12.9306
F12 | 12.8898
G01 | 20.2741
G02 | 21.5978
G03 | 19.8264
G04 | 20.1345
G05 | 20.6049
G06 | 21.9131
G07 | 11.6583
G08 | 11.6704
G09 | 11.688
G10 | 15.1786
G11 | 15.2167
G12 | 15.1839
H01 | 21.9741
H02 | 23.7367
H03 | 21.7893
H04 | 22.3819
H05 | 22.193
H06 | 24.2687
H07 | 11.9743
H08 | 11.969
H09 | 11.9911
H10 | 15.4462
H11 | 15.5403
H12 | 15.5313

## Sample library dilution
Well1 | Well2 | Well3 | Sample name | Dilution Factor | X | Relative amount/standard amount
A01 | A02 | A03 | std1 | 10 | -1 | 0.1
A04 | A05 | A06 | std2 | 100 | -2 | 0.01
A07 | A08 | A09 | std3 | 1000 | -3 | 0.001
A10 | A11 | A12 | std4 | 10000 | -4 | 0.0001
B01 | B02 | B03 | std5 | 100000 | -5 | 1e-05
C07 | C08 | C09 | sample library 1 dilution 1 | 2000
C10 | C11 | C12 | sample library 1 dilution 2 | 20000
D07 | D08 | D09 | sample library 2 dilution 1 | 2000
D10 | D11 | D12 | sample library 2 dilution 2 | 20000
E07 | E08 | E09 | sample library 3 dilution 1 | 2000
E10 | E11 | E12 | sample library 3 dilution 2 | 20000
F07 | F08 | F09 | sample library 4 dilution 1 | 2000
F10 | F11 | F12 | sample library 4 dilution 2 | 20000
G07 | G08 | G09 | sample library 5 dilution 1 | 2000
G10 | G11 | G12 | sample library 5 dilution 2 | 20000
H07 | H08 | H09 | sample library 6 dilution 1 | 2000
H10 | H11 | H12 | sample library 6 dilution 2 | 20000
C01 |  |  | library1 Spike1 | 2000
C02 |  |  | library1 Spike2 | 2000
C03 |  |  | library1 Spike3 | 2000
C04 |  |  | library1 Spike4 | 2000
D01 |  |  | library2 Spike1 | 2000
D02 |  |  | library2 Spike2 | 2000
D03 |  |  | library2 Spike3 | 2000
D04 |  |  | library2 Spike4 | 2000
E01 |  |  | library3 Spike1 | 2000
E02 |  |  | library3 Spike2 | 2000
E03 |  |  | library3 Spike3 | 2000
E04 |  |  | library3 Spike4 | 2000
F01 |  |  | library4 Spike1 | 2000
F02 |  |  | library4 Spike2 | 2000
F03 |  |  | library4 Spike3 | 2000
F04 |  |  | library4 Spike4 | 2000
G01 |  |  | library5 Spike1 | 2000
G02 |  |  | library5Spike2 | 2000
G03 |  |  | library5 Spike3 | 2000
G04 |  |  | library5 Spike4 | 2000
H01 |  |  | library6 Spike1 | 2000
H02 |  |  | library6 Spike2 | 2000
H03 |  |  | library6 Spike3 | 2000
H04 |  |  | library6 Spike4 | 2000
B04 | B05 | B06 | NTC
B07 | B08 | B09 | NTC
B10 | B11 | B12 | NTC

## Library information
Library # | Well1 | Well2 | Well3 | Well4 | Total primer pair | Library size | Table 1
sample library 1 | C01 | C02 | C03 | C04 | 6943 | 220 | Cat # | Panel name | # primer pairs
sample library 2 | D01 | D02 | D03 | D04 | 6943 | 220 | NGHS-001X | Breast cancer | 2915
sample library 3 | E01 | E02 | E03 | E04 | 6943 | 220 | NGHS-002X | Colorectal cancer | 1954
sample library 4 | F01 | F02 | F03 | F04 | 6943 | 220 | NGHS-003X | Myeloid Neoplasms | 2536
sample library 5 | G01 | G02 | G03 | G04 | 6943 | 220 | NGHS-004X | Liver cancer | 2052
sample library 6 | H01 | H02 | H03 | H04 | 6943 | 220 | NGHS-005X | Lung cancer | 3586
[7 empty] | NGHS-006X | Ovarian cancer | 2021
[7 empty] | NGHS-007X | Prostate cancer | 1837
[7 empty] | NGHS-008X | Gastric cancer | 2377
[7 empty] | NGHS-009X | Cardiomyopathy | 2657
[7 empty] | NGHS-011X | Carrier testing | 6943
[7 empty] | NGHS-013X | Cancer predisposition | 6582
[7 empty] | NGHS-501X | Comprehensive cancer | 7951
[7 empty] | NGHS-101X | Clinically relevant | 602
[7 empty] | NGHS-102X | BRCA1 and BRCA2 | 250
[7 empty] | NGHS-201X | Actionable mutations | 118

## QC report
NTC QC | [5 empty] | Well1 | Well2 | Well3 | Sample name | Dilution Fold | Value1 | Value2 | Value3 | Ave CT (Y) | STDEV
[6 empty] | A01 | A02 | A03 | std1 | 10 | 9.719 | 9.7991 | 9.7376 | 9.7519 | 0.041924
Well1 | Well2 | Well3 | Sample name | Average | Pass/Fail | A04 | A05 | A06 | std2 | 100 | 13.3176 | 13.2803 | 13.3032 | 13.3004 | 0.01878
B04 | B05 | B06 | NTC | 40 | Pass | A07 | A08 | A09 | std3 | 1000 | 16.7364 | 16.7909 | 16.8021 | 16.7765 | 0.035141
B07 | B08 | B09 | NTC | 40 | Pass | A10 | A11 | A12 | std4 | 10000 | 20.333 | 20.4496 | 20.3468 | 20.3765 | 0.063727
B10 | B11 | B12 | NTC | 40 | Pass | B01 | B02 | B03 | std5 | 100000 | 23.8822 | 23.8979 | 23.8454 | 23.8751 | 0.026952
[6 empty] | C07 | C08 | C09 | sample library 1 dilution 1 | 2000 | 11.9539 | 11.967 | 11.9829 | 11.9679 | 0.014488
[6 empty] | C10 | C11 | C12 | sample library 1 dilution 2 | 20000 | 15.4695 | 15.4925 | 15.5201 | 15.494 | 0.025346
[6 empty] | D07 | D08 | D09 | sample library 2 dilution 1 | 2000 | 8.6829 | 8.6929 | 8.7377 | 8.7045 | 0.029183
Slope QC | [5 empty] | D10 | D11 | D12 | sample library 2 dilution 2 | 20000 | 13.3192 | 13.3044 | 13.2823 | 13.302 | 0.018593
[6 empty] | E07 | E08 | E09 | sample library 3 dilution 1 | 2000 | 4.8088 | 4.8192 | 4.8089 | 4.8123 | 0.0059763
| slope | intercept | RSQ | Efficiency | Pass/Fail | E10 | E11 | E12 | sample library 3 dilution 2 | 20000 | 8.1884 | 8.2265 | 8.1931 | 8.2027 | 0.020781
| -3.5323 | 6.2193 | 0.99998 | 0.91913 | Pass | F07 | F08 | F09 | sample library 4 dilution 1 | 2000 | 9.3232 | 9.3416 | 9.3502 | 9.3383 | 0.013806
[6 empty] | F10 | F11 | F12 | sample library 4 dilution 2 | 20000 | 12.9285 | 12.9306 | 12.8898 | 12.9163 | 0.022971
[6 empty] | G07 | G08 | G09 | sample library 5 dilution 1 | 2000 | 11.6583 | 11.6704 | 11.688 | 11.6722 | 0.014963
[6 empty] | G10 | G11 | G12 | sample library 5 dilution 2 | 20000 | 15.1786 | 15.2167 | 15.1839 | 15.1931 | 0.020665
[6 empty] | H07 | H08 | H09 | sample library 6 dilution 1 | 2000 | 11.9743 | 11.969 | 11.9911 | 11.9781 | 0.011523
[6 empty] | H10 | H11 | H12 | sample library 6 dilution 2 | 20000 | 15.4462 | 15.5403 | 15.5313 | 15.506 | 0.051918
[6 empty] | B04 | B05 | B06 | NTC |  | 40 | 40 | 40 | 40 | 0
[6 empty] | B07 | B08 | B09 | NTC |  | 40 | 40 | 40 | 40 | 0
[6 empty] | B10 | B11 | B12 | NTC |  | 40 | 40 | 40 | 40 | 0

## Results
Well1 | Well2 | Well3 | Sample name | pM | Template dilution factor | Av CT (Y)
A01 | A02 | A03 | std1 |  |  | 9.7519
A04 | A05 | A06 | std2 |  |  | 13.3004
A07 | A08 | A09 | std3 |  |  | 16.7765
A10 | A11 | A12 | std4 |  |  | 20.3765
B01 | B02 | B03 | std5 |  |  | 23.8751
C07 | C08 | C09 | sample library 1 dilution 1 | 3258.1145 | 814.5286 | 11.9679
C10 | C11 | C12 | sample library 1 dilution 2 | 3271.1789 | 817.7947 | 15.494
D07 | D08 | D09 | sample library 2 dilution 1 | 27344.563 | 6836.1407 | 8.7045
D10 | D11 | D12 | sample library 2 dilution 2 | 13654.8523 | 3425.0921 | 13.302
E07 | E08 | E09 | sample library 3 dilution 1 | 345747.9256 | 86436.9814 | 4.8123
E10 | E11 | E12 | sample library 3 dilution 2 | 379254.9724 | 94813.7431 | 8.2027
F07 | F08 | F09 | sample library 4 dilution 1 | 18089.2021 | 4522.3005 | 9.3383
F10 | F11 | F12 | sample library 4 dilution 2 | 17558.0756 | 4389.5189 | 12.9163
G07 | G08 | G09 | sample library 5 dilution 1 | 3950.7197 | 987.6799 | 11.6722
G10 | G11 | G12 | sample library 5 dilution 2 | 3980.2888 | 995.0722 | 15.1931
H07 | H08 | H09 | sample library 6 dilution 1 | 3236.5251 | 809.1313 | 11.9781
H10 | H11 | H12 | sample library 6 dilution 2 | 3245.8882 | 811.472 | 15.506 | Notes:
If both dilutions result in CT values within the dynamic range of standards, use pM values from the most concentrated library dilution. If one dilution only falls within the dynamic range of standards, choose the pM value from the dilution that yields a CT value within the dynamic range of standards.
If CT values from both dilutions fall out of range (low CT), the library is too concentrated. We recommend checking the concentration using BioAnalyzer or QIAxcel and diluting the library to about 1 nM, and repeating library quantification using the GeneRead Library Quant System.
If CT values from both dilutions fall out of range (high CT), the library concentration is too low and not good for proceeding to sequencing. Library construction has to be repeated.
Target Enrichment QC
Library # | QC Score | QC Result
sample library 1 | 3.2383 | Pass
sample library 2 | 1.5536 | Pass
sample library 3 | 30.6952 | Fail
sample library 4 | 21.892 | Fail
sample library 5 | 12.1938 | Fail
sample library 6 | 4.1116 | Pass

## Calculation
Well | raw CT | Well1 | Well2 | Well3 | Sample name | Dilution fold | Value1 | Value2 | Value3 | Av CT (Y) | STDEV | X | Relative amount/standard amount | Copies/ml | Template dilution factor (dilution fold from original library) | pM | slope | intercept | RSQ | Efficiency
A01 | 9.719 | A01 | A02 | A03 | std1 | 10 | 9.719 | 9.7991 | 9.7376 | 9.7519 | 0.041924 | -1 |  |  |  |  | -3.5323 | 6.2193 | 0.99998 | 0.91913
A02 | 9.7991 | A04 | A05 | A06 | std2 | 100 | 13.3176 | 13.2803 | 13.3032 | 13.3004 | 0.01878 | -2
A03 | 9.7376 | A07 | A08 | A09 | std3 | 1000 | 16.7364 | 16.7909 | 16.8021 | 16.7765 | 0.035141 | -3 |  |  |  |  | Library # | Well1 | Well2 | Well3 | Well4 | Total Primer Pairs | Value1 | Value2 | Value3 | Value4 | Total | Experimental | Theoretical | QC Score | QC Result
A04 | 13.3176 | A10 | A11 | A12 | std4 | 10000 | 20.333 | 20.4496 | 20.3468 | 20.3765 | 0.063727 | -4 |  |  |  |  | sample library 1 | C01 | C02 | C03 | C04 | 6943 | 8.7012e-07 | 3.5721e-07 | 3.8035e-06 | 3.3927e-06 | 8.4235e-06 | 0.00035581 | 0.0011522 | 3.2383 | Pass
A05 | 13.2803 | B01 | B02 | B03 | std5 | 100000 | 23.8822 | 23.8979 | 23.8454 | 23.8751 | 0.026952 | -5 |  |  |  |  | sample library 2 | D01 | D02 | D03 | D04 | 6943 | 3.4479e-05 | 1.5405e-05 | 3.379e-05 | 9.3232e-05 | 0.00017691 | 0.0017901 | 0.0011522 | 1.5536 | Pass
A06 | 13.3032 | C07 | C08 | C09 | sample library 1 dilution 1 | 2000 | 11.9539 | 11.967 | 11.9829 | 11.9679 | 0.014488 | -1.6275 | 0.023579 | 1961384944.7487 | 814.5286 | 3258.1145 | sample library 3 | E01 | E02 | E03 | E04 | 6943 | 0.026125 | 0.017089 | 0.026678 | 0.027184 | 0.097076 | 0.035368 | 0.0011522 | 30.6952 | Fail
A07 | 16.7364 | C10 | C11 | C12 | sample library 1 dilution 2 | 20000 | 15.4695 | 15.4925 | 15.5201 | 15.494 | 0.025346 | -2.6257 | 0.0023674 | 1969249669.5452 | 817.7947 | 3271.1789 | sample library 4 | F01 | F02 | F03 | F04 | 6943 | 0.0010102 | 0.00046135 | 0.00092764 | 0.00080609 | 0.0032053 | 0.025225 | 0.0011522 | 21.892 | Fail
A08 | 16.7909 | D07 | D08 | D09 | sample library 2 dilution 1 | 2000 | 8.6829 | 8.6929 | 8.7377 | 8.7045 | 0.029183 | -0.70356 | 0.1979 | 16461426906.1038 | 6836.1407 | 27344.563 | sample library 5 | G01 | G02 | G03 | G04 | 6943 | 0.00010496 | 4.4287e-05 | 0.00014053 | 0.00011496 | 0.00040473 | 0.01405 | 0.0011522 | 12.1938 | Fail
A09 | 16.8021 | D10 | D11 | D12 | sample library 2 dilution 2 | 20000 | 13.3192 | 13.3044 | 13.2823 | 13.302 | 0.018593 | -2.0051 | 0.0098822 | 8220221059.1754 | 3425.0921 | 13654.8523 | sample library 6 | H01 | H02 | H03 | H04 | 6943 | 3.4653e-05 | 1.0983e-05 | 3.909e-05 | 2.6564e-05 | 0.00011129 | 0.0047376 | 0.0011522 | 4.1116 | Pass
A10 | 20.333 | E07 | E08 | E09 | sample library 3 dilution 1 | 2000 | 4.8088 | 4.8192 | 4.8089 | 4.8123 | 0.0059763 | 0.39833 | 2.5022 | 208140251230.6917 | 86436.9814 | 345747.9256
A11 | 20.4496 | E10 | E11 | E12 | sample library 3 dilution 2 | 20000 | 8.1884 | 8.2265 | 8.1931 | 8.2027 | 0.020781 | -0.5615 | 0.27447 | 228311493383.9757 | 94813.7431 | 379254.9724
A12 | 20.3468 | F07 | F08 | F09 | sample library 4 dilution 1 | 2000 | 9.3232 | 9.3416 | 9.3502 | 9.3383 | 0.013806 | -0.88301 | 0.13091 | 10889699638.3811 | 4522.3005 | 18089.2021
B01 | 23.8822 | F10 | F11 | F12 | sample library 4 dilution 2 | 20000 | 12.9285 | 12.9306 | 12.8898 | 12.9163 | 0.022971 | -1.896 | 0.012707 | 10569961533.5433 | 4389.5189 | 17558.0756
B02 | 23.8979 | G07 | G08 | G09 | sample library 5 dilution 1 | 2000 | 11.6583 | 11.6704 | 11.688 | 11.6722 | 0.014963 | -1.5438 | 0.028592 | 2378333251.8054 | 987.6799 | 3950.7197
B03 | 23.8454 | G10 | G11 | G12 | sample library 5 dilution 2 | 20000 | 15.1786 | 15.2167 | 15.1839 | 15.1931 | 0.020665 | -2.5405 | 0.0028806 | 2396133860.3806 | 995.0722 | 3980.2888
B04 | 40 | H07 | H08 | H09 | sample library 6 dilution 1 | 2000 | 11.9743 | 11.969 | 11.9911 | 11.9781 | 0.011523 | -1.6304 | 0.023423 | 1948388124.9845 | 809.1313 | 3236.5251
B05 | 40 | H10 | H11 | H12 | sample library 6 dilution 2 | 20000 | 15.4462 | 15.5403 | 15.5313 | 15.506 | 0.051918 | -2.6291 | 0.0023491 | 1954024684.3435 | 811.472 | 3245.8882
B06 | 40 | C01 |  |  | library1 Spike1 | 2000 | 27.6263 |  |  | 27.6263 |  | -6.0604 | 8.7012e-07
B07 | 40 | C02 |  |  | library1 Spike2 | 2000 | 28.992 |  |  | 28.992 |  | -6.4471 | 3.5721e-07
B08 | 40 | C03 |  |  | library1 Spike3 | 2000 | 25.3635 |  |  | 25.3635 |  | -5.4198 | 3.8035e-06
B09 | 40 | C04 |  |  | library1 Spike4 | 2000 | 25.5388 |  |  | 25.5388 |  | -5.4695 | 3.3927e-06
B10 | 40 | D01 |  |  | library2 Spike1 | 2000 | 21.9818 |  |  | 21.9818 |  | -4.4624 | 3.4479e-05
B11 | 40 | D02 |  |  | library2 Spike2 | 2000 | 23.2177 |  |  | 23.2177 |  | -4.8123 | 1.5405e-05
B12 | 40 | D03 |  |  | library2 Spike3 | 2000 | 22.0128 |  |  | 22.0128 |  | -4.4712 | 3.379e-05
C01 | 27.6263 | D04 |  |  | library2 Spike4 | 2000 | 20.4558 |  |  | 20.4558 |  | -4.0304 | 9.3232e-05
C02 | 28.992 | E01 |  |  | library3 Spike1 | 2000 | 11.8106 |  |  | 11.8106 |  | -1.5829 | 0.026125
C03 | 25.3635 | E02 |  |  | library3 Spike2 | 2000 | 12.4618 |  |  | 12.4618 |  | -1.7673 | 0.017089
C04 | 25.5388 | E03 |  |  | library3 Spike3 | 2000 | 11.7785 |  |  | 11.7785 |  | -1.5739 | 0.026678
C05 | 28.8413 | E04 |  |  | library3 Spike4 | 2000 | 11.7497 |  |  | 11.7497 |  | -1.5657 | 0.027184
C06 | 29.2539 | F01 |  |  | library4 Spike1 | 2000 | 16.8004 |  |  | 16.8004 |  | -2.9956 | 0.0010102
C07 | 11.9539 | F02 |  |  | library4 Spike2 | 2000 | 18.0028 |  |  | 18.0028 |  | -3.336 | 0.00046135
C08 | 11.967 | F03 |  |  | library4 Spike3 | 2000 | 16.9313 |  |  | 16.9313 |  | -3.0326 | 0.00092764
C09 | 11.9829 | F04 |  |  | library4 Spike4 | 2000 | 17.1467 |  |  | 17.1467 |  | -3.0936 | 0.00080609
C10 | 15.4695 | G01 |  |  | library5 Spike1 | 2000 | 20.2741 |  |  | 20.2741 |  | -3.979 | 0.00010496
C11 | 15.4925 | G02 |  |  | library5Spike2 | 2000 | 21.5978 |  |  | 21.5978 |  | -4.3537 | 4.4287e-05
C12 | 15.5201 | G03 |  |  | library5 Spike3 | 2000 | 19.8264 |  |  | 19.8264 |  | -3.8522 | 0.00014053
D01 | 21.9818 | G04 |  |  | library5 Spike4 | 2000 | 20.1345 |  |  | 20.1345 |  | -3.9395 | 0.00011496
D02 | 23.2177 | H01 |  |  | library6 Spike1 | 2000 | 21.9741 |  |  | 21.9741 |  | -4.4603 | 3.4653e-05
D03 | 22.0128 | H02 |  |  | library6 Spike2 | 2000 | 23.7367 |  |  | 23.7367 |  | -4.9593 | 1.0983e-05
D04 | 20.4558 | H03 |  |  | library6 Spike3 | 2000 | 21.7893 |  |  | 21.7893 |  | -4.4079 | 3.909e-05
D05 | 22.4246 | H04 |  |  | library6 Spike4 | 2000 | 22.3819 |  |  | 22.3819 |  | -4.5757 | 2.6564e-05
D06 | 23.3641 | B04 | B05 | B06 | NTC |  | 40 | 40 | 40 | 40 | 0
D07 | 8.6829 | B07 | B08 | B09 | NTC |  | 40 | 40 | 40 | 40 | 0
D08 | 8.6929 | B10 | B11 | B12 | NTC |  | 40 | 40 | 40 | 40 | 0
D09 | 8.7377
D10 | 13.3192
D11 | 13.3044
D12 | 13.2823
E01 | 11.8106
E02 | 12.4618
E03 | 11.7785
E04 | 11.7497
E05 | 12.7409
E06 | 13.156
E07 | 4.8088
E08 | 4.8192
E09 | 4.8089
E10 | 8.1884
E11 | 8.2265
E12 | 8.1931
F01 | 16.8004
F02 | 18.0028
F03 | 16.9313
F04 | 17.1467
F05 | 17.3042
F06 | 18.5281
F07 | 9.3232
F08 | 9.3416
F09 | 9.3502
F10 | 12.9285
F11 | 12.9306
F12 | 12.8898
G01 | 20.2741
G02 | 21.5978
G03 | 19.8264
G04 | 20.1345
G05 | 20.6049
G06 | 21.9131
G07 | 11.6583
G08 | 11.6704
G09 | 11.688
G10 | 15.1786
G11 | 15.2167
G12 | 15.1839
H01 | 21.9741
H02 | 23.7367
H03 | 21.7893
H04 | 22.3819
H05 | 22.193
H06 | 24.2687
H07 | 11.9743
H08 | 11.969
H09 | 11.9911
H10 | 15.4462
H11 | 15.5403
H12 | 15.5313